## Design Steps
Test name | Test ID | Test External ID | Step ID | Step Name | Step description | Step Parameters | Expected | Step type | Has Attachments
IQ Sustainability Test |  |  |  | Step0 | Run this TP when new detector design phase in(new structure or new material) |  | NA | NA
IQ Sustainability Test | *** | *** | *** | Step1 | Phantoms used in this test procedure: kyoto head phantom Xiaomei phantom 12.5cm Water Phantom 20cm Water Phantom | --- | Information only. No expected results. | Setup | add a kyoto head phantom & xiaomei phantom picture
IQ Sustainability Test | *** | *** | *** | Step2 | Confirmed system already completed detailed cal, fast cal. | --- | Information only. No expected results. | Setup | No
IQ Sustainability Test | *** | *** | *** | Step3 | Open a shell, enter "tubetemp -t" to check tube temperature. Make sure the temperature within 200~300. | --- | Information only. No expected results. | Setup | No
IQ Sustainability Test | *** | *** | *** | Step4 | Mount and level the kyoto head phantom on the head holder. Center the phantom in the Z direction and roughly center the phantom in the X and Y directions. | --- | The phantom is mounted, leveled, and centered in the Z direction. | Setup | No
IQ Sustainability Test | *** | *** | *** | Step5 | Set up a new exam with the following scan parameters. Series1 Start: I37.5/End: S37.5/SFOV: Large Body/DFOV: 35.0 cm/kV: 120/mA: 380/Scan Type: Axial/Hi Res/Mode: Off/Rotation Time: 1.0s/Detector Coverage: 80 mm/Thickness: 5.0mm/Recon Type: Stnd/WW: 400/WL: 40 Set a landmark using the internal laser lights | --- | Series was set up with the specified parameters. | Setup | No
IQ Sustainability Test | *** | *** | *** | Step6 | Acquire the scan data for the series. Record the exam number. Continue and close the exam. | --- | Record exam number and series number. | Datapoint | No
IQ Sustainability Test | *** | *** | *** | Step7 | Mount and level the xiaomei head phantom on the head holder. Center the phantom in the Z direction and roughly center the phantom in the X and Y directions. | --- | The phantom is mounted, leveled, and centered in the Z direction.
IQ Sustainability Test | *** | *** | *** | Step8 | Set up a new exam with the following scan parameters. Series1 Start: I37.5/End: S37.5/SFOV: Large Body/DFOV: 35.0 cm/kV: 120/mA: 380/Scan Type: Axial/Hi Res/Mode: Off/Rotation Time: 1.0s/Detector Coverage: 80 mm/Thickness: 5.0mm/Recon Type: Stnd/WW: 400/WL: 40 Set a landmark using the internal laser lights | --- | Series 1 was set up with the specified parameters. | Setup | No
IQ Sustainability Test | *** | *** | *** | Step9 | Acquire the scan data for the series. Record the exam number. Continue and close the exam. | --- | Record exam number and series number. | Setup | No
IQ Sustainability Test | *** | *** | *** | Step10 | Mount and level the 20cm water phantom on the phantom holder. Roughly center the phantom in the Z direction and using the Center Phantom Iso tool from the service tool to center the phantom in the X and Y directions. | --- | The phantom is mounted, leveled, and centered in the Z direction. | Datapoint | No
IQ Sustainability Test | *** | *** | *** | Step11 | Set up a new exam with the following scan parameters. Series1 Start: I37.5/End: S37.5/SFOV: Large Body/DFOV: 35.0 cm/kV: 120/mA: 380/Scan Type: Axial/Hi Res Mode: Off/Rotation Time: 1.0s/Detector Coverage: 80 mm/Thickness: 5.0mm/Recon Type: Stnd/WW: 40/WL: 0 Series2 Start: I37.5/End: S37.5/SFOV: Large Body/DFOV: 35.0 cm/kV: 120/mA: 380/Scan Type: Axial/Hi Res Mode: Off/Rotation Time: 0.5s/Detector Coverage: 80 mm/Thickness: 5.0mm/Recon Type: Stnd/WW: 40/WL: 0 Series3 Start: I37.5/End: S37.5/SFOV: Large Body/DFOV: 35.0 cm/kV: 120/mA: 380/Scan Type: Helical/Hi Res Mode: Off/Rotation Time: 0.5s/Detector Coverage: 40 mm/Pitch: 0.984:1/Thickness: 5.0mm/Recon Type: Stnd/WW: 40/WL: 0 Series4 Start: I37.5/End: S37.5/SFOV: Large Body/DFOV: 35.0 cm/kV mode: GSI/kV: 80 - 140/mA: 380/Scan Type: Helical/Hi Res Mode: Off/Rotation Time: 0.5s/Detector Coverage: 40 mm/Pitch: 0.516:1/Thickness: 5.0mm/Recon Type: Stnd/WW: 40/WL: 0 Set a landmark using the internal laser lights | --- | Series 1 was set up with the specified parameters. | Checkpoint | No
IQ Sustainability Test | *** | *** | *** | Step12 | Acquire the scan data for the series. Record the exam number. Continue and close the exam. | --- | Record exam number and series number. | Setup | No
IQ Sustainability Test | *** | *** | *** | Step13 | Mount and level the 12.5cm water phantom on the phantom holder. Roughly center the phantom in the Z direction and using the Center Phantom Iso tool from the service tool to center the phantom in the X and Y directions.
IQ Sustainability Test | *** | *** | *** | Step14 | Set up a new exam with the following scan parameters. Series1 Start: I37.5/End: S37.5/SFOV: Large Body/DFOV: 35.0 cm/kV: 120/mA: 380/Scan Type: Axial/Hi Res Mode: Off/Rotation Time: 1.0s/Detector Coverage: 80 mm/Thickness: 5.0mm/Recon Type: Stnd/WW: 40/WL: 0 Set a landmark using the internal laser lights | --- | Series 1 was set up with the specified parameters. | Setup | No
IQ Sustainability Test | *** | *** | *** | Step15 | Acquire the scan data for the series. Record the exam number. Continue and close the exam. | --- | Record exam number and series number. | Datapoint | No
IQ Sustainability Test | *** | *** | *** | Step16 | Visually check all images of the step 6, 9, 12, 15 for artifacts. The results is baseline. | --- | No obvious artifacts on all images. | Checkpoint | No
IQ Sustainability Test | *** | *** | *** | Step17 | Leave it overnight. In next morning, checked tube temperature, make sure it is less than 60 'C. | --- | Information only. No expected results. | Setup | No
IQ Sustainability Test | *** | *** | *** | Step18 | Repeat step 4~16. | --- | No obvious degradation compred with baseline. | Checkpoint | No
IQ Sustainability Test | *** | *** | *** | Step19 | Leave it overnight. In next morning, checked tube temperature, make sure it is less than 60 'C. | --- | Information only. No expected results. | Setup | No
IQ Sustainability Test | *** | *** | *** | Step20 | Repeat step 4~16. | --- | No obvious degradation compred with baseline. | Checkpoint | No
IQ Sustainability Test | *** | *** | *** | Step21 | Leave it overnight. In next morning, checked tube temperature, make sure it is less than 60 'C. |  | Information only. No expected results. | Setup | No
IQ Sustainability Test | *** | *** | *** | Step22 | Repeat step 4~16. |  | No obvious degradation compred with baseline. | Checkpoint | No